Radiation Safety Report: Room Design for New PET 2019, Rambam

Date: 	2019.10.03

John A. Kennedy, PhD.

Chief Physicist

Dept. of Nuclear Medicine

Rambam Health Care Campus

P.O.B. 9602, Haifa, Israel, 3109601

Phone: 972-4-777-2815

Fax: 972-4-777-3011

j_kennedy@rambam.health.gov.il

Introduction

These are Rambam Health Care Campus estimates of the necessary shielding for the PET and uptake rooms. Calculations were made in accordance with Mark T. Madsen, et al., AAPM Task Group 108: PET and PET/CT Shielding Requirements, Med. Phys. 33 (1):4, 2006. The contribution of the CT the dose is minimal because the scans are low to moderate in current and voltage ( ≤ 90 mAs, ≤ 140 kV), limited in frequency ( ≤ 3 scans /hour), and shielded by a minimum of 20 cm of concrete or 2 mm lead (or 4 mm lead equivalent glass).  Figure 1 gives the details of the room shielding.

These are shielding calculations. All other radiation safety regulations must be applied to the rooms such as PVC flooring, non-absorbent walls, and sinks in the PET uptake room, PET WC (toilet), and PET/CT exam room.

Alternative plans are possible as long as calculations are provided to show that they meet regulations. Also, for 511 keV photons, lead (Pb) can be substituted for poured concrete (2.35 g/cm3 density) shielding according to calculations based on the above reference. Shielding using barium, Ba, (e.g. barium sulfide drywall) is NOT appropriate for 511 keV photons. A partial table of equivalence is:

| Concrete (2.35 g/cm3) and Lead (Pb) Equivalent Thicknesses |  |  |
| --- | --- | --- |
| Concrete Thickness (cm) | Lead Thickness (mm) | Transmission Factor (at 511 keV) |
| 5 | 2.7 | 0.72179 |
| 10 | 6.9 | 0.39867 |
| 15 | 11.6 | 0.19319 |
| 20 | 16.6 | 0.09036 |
| 25 | 21.6 | 0.04194 |
| 30 | 26.6 | 0.01944 |
| 35 | 31.5 | 0.00901 |
| 40 | 36.5 | 0.00417 |

The report consists of the following parts:

- Parameters used for shielding calculations.

- Equations used for shielding calculations.

- List of PET/CT radiation shielding.

- Figure 1: floor plan with reference codes for radiation shielding.

- Sample calculation: Exposure to the hematology wing above the uptake room.

- Discussion on Calculations found in Tables

- Conclusion

- Tables

| Parameters used in shielding calculations. |  |  |  |
| --- | --- | --- | --- |
| Parameter |  | Value | Units |
| Radioisotope |  | F-18 FDG |  |
| Half-life | T1/2 | 109.77 | minutes |
| Patient gamma factor | Γ | 0.092 | μSv·m²/MBq/h |
| Lead (Pb) Archer Eq. values: |  |  |  |
| alpha | α | 1.543 | cm−1 |
| beta | β | −0.4408 | cm−1 |
| gamma | γ | 2.136 |  |
| Concrete (2.35 g/cm3) Archer Eq values: |  |  |  |
| alpha | α | 0.1539 | cm−1 |
| beta | β | −0.1161 | cm−1 |
| gamma | γ | 2.0752 |  |
| Injected activity | A0 | 10 | mCi |
|  |  | 370 | MBq |
| Uptake period | tU | 60 | minutes |
| At injection location | tU1 | 5 | minutes |
| At waiting location | tU2 | 50 | minutes |
| At WC (toilet) | tU3 | 5 | minutes |
| Imaging period | tI | 20 | minutes |
| Voiding within uptake period | FV | 15 | % |
| Shielding by scanner | FS | 20 | % |
| Dose reduction factors: |  |  |  |
| During uptake | RtU | 0.832383 |  |
| During uptake at injection location | RtU1 | 0.984378 |  |
| During uptake at waiting location | RtU2 | 0.857518 |  |
| During uptake at WC | RtU3 | 0.984378 |  |
| During imaging | RtI | 0.939431 |  |
| Uptake decay time factor after injection | FU1 | 0.968921 |  |
| Uptake decay time factor after waiting | FU2 | 0.706593 |  |
| Uptake decay time factor after WC | FU3 | 0.684633 |  |
| Number of patients per week | Nw | 160 | week−1 |
| Yearly Dose Limit: |  |  |  |
| Uncontrolled area |  | 0.2 | mSv/year |
| Controlled area |  | 2 | mSv/year |
| Weekly Dose limit: | P |  |  |
| Uncontrolled area |  | 4 | μSv/week |
| Controlled |  | 40 | μSv/week |
| Hourly Dose Limit: |  |  |  |
| Public |  | 20 | μSv/h |
|  |  | 2 | mrem/h |
| Occupancy Factor: | T |  |  |
| Offices/workrooms/patient rooms |  | 1 |  |
| Corridors |  | 0.25 |  |
| Outside/toilets/stairways/janitor closets |  | 0.0625 |  |
| (Continued on next page.) |  |  |  |
| Parameters used in shielding calculations (continued from previous page). |  |  |  |
| Parameter |  | Value | Units |
| Distances: | d |  |  |
| Person to wall |  | 0.3 | m |
| Prone person to floor |  | 1.0 | m |
| Sitting person to floor |  | 0.5 | m |

Equations Used for Shielding Calculations

Archer Equation for determining broad beam transmission factors:

B = {(1+(β/α))eαγx−(β/α)}(−1/γ)

Archer Equation for determining shielding thicknesses from required transmission factors.

x = (1/αγ)ln{[B−γ+(β/α)]/[1+(β/α)]}

Dose rate reduction factors:

Rt=[1/ln(2)](T1/2/t){1-exp[-ln(2)t/T1/2]}

Dose at point d meters from patient during uptake time:

D = ΓA0tURtU/d2

Weekly dose is DNw. Weekly dose limit is P and occupancy factor is T. Therefore, the required transmission factor is

B = P/(TNwD)

Uptake time decay factor:

FU = exp[-ln(2)tU/T1/2]

Dose at point d meters from patient during imaging time:

D = ΓA0tIRtIFU(1-FV)(1-FS)/d2

Dose (Ď) at distance of 1 m from uptake patient:

Ď = ΓA0tURtU

Dose (Ď) at distance of 1 m from patient during imaging:

Ď = ΓA0tIRtIFU(1-FV)(1-FS)

Radiation field or dose rate (R) at 1 m from uptake patient:

R = Ď/tU

R = ΓA0RtU

Radiation field or dose rate (R) at 1 m from patient during imaging:

R = Ď/tI

R = ΓA0RtIFU(1-FV)(1-FS)

List of New 2019 PET/CT Rooms’ Radiation Shielding Requirements (See Figure 1)

| Code | Description | Shielding | Comment |
| --- | --- | --- | --- |
| W1 | Wall PET Uptake Room N | 16 cm concrete** | Hallway |
| W2 | Wall PET Uptake Room W | 36 cm concrete | Shields SPECT control room |
| W3 | Wall PET Uptake Room S | 31 cm concrete | Uncontrolled area south |
| W4 | Wall PET Uptake Room E | 32 cm concrete | Uncontrolled area east |
| W5 | Wall PET WC S | 20 cm concrete | Uncontrolled area south |
| W6 | Wall PET WC E | 20 cm concrete | Uncontrolled area east |
| W7 | Wall PET WC N | 27 cm concrete | Uncontrolled area north |
| W8 | Wall PET/CT Exam Room S | 2 mm Pb | Entire southern wall |
| W9 | Wall PET/CT Exam Room E | 25 cm concrete | Uncontrolled area east |
| W10 | Wall PET/CT Exam Room N | 20 cm concrete | Uncontrolled area north |
| W11 | Wall PET/CT Exam Room W | 20 cm concrete | Shields PET control room |
| H1 | Window PET/CT Control Room | 4 mm Pb equivalent | 120 cm x 80 cm |
| D1 | Door PET Uptake Room | 4 mm Pb | Sliding - automatic |
| D2 | Door PET/CT Exam Room | 2 mm Pb | Sliding - automatic |
| D3 | Door PET/CT Control Room | 2 mm Pb | Wing |
| D4 | PET WC Door | none | Wing |
| C1 | Ceiling PET Uptake Room | 34 cm concrete | Uncontrolled area above |
| C2 | Ceiling PET WC | 17 cm concrete | Uncontrolled area above |
| C3 | Ceiling PET/CT Exam Room | 25 cm concrete | Uncontrolled area above |
| *N,W,S,E = north, west, south, east walls **All concrete should be poured (בֶּטוֹן יָצוּק) and of normal density (minimum 2.35 g/cc). Additional shielding is provided to the technician doing the injection by standard 20 mm Pb shield used routinely at Rambam since 2012. |  |  |  |

List of New 2019 PET/CT Rooms’ Sources and Targets (See Figure 1)

| Code | Description | Comment |
| --- | --- | --- |
| S1 | Source | Patient being injected for 5 minutes |
| S2 | Source | Patient uptake location for 50 minutes: one location of three |
| S3 | Source | Patient uptake location for 50 minutes: one location of three |
| S4 | Source | Patient uptake location for 50 minutes: one location of three |
| S5 | Source | Patient WC location for 5 minutes |
| S6 | Source | Patient scanned for 20 minutes |
| T1 | Target | Technician performing injection* |
| T2 | Target | SPECT/CT control room |
| T3 | Target | Outside in garden behind a structural pillar (uncontrolled)** |
| T4 | Target | Stairwell to Internal Medicine H (uncontrolled) |
| T5 | Target | Janitor closet (uncontrolled) |
| T6 | Target | WC (uncontrolled) |
| T7 | Target | WC (uncontrolled) |
| T8 | Target | Hallway (uncontrolled) |
| T9 | Target | Hallway outside PET uptake room wall |
| T10 | Target | Hallway outside PET/CT exam room wall |
| T11 | Target | Hallway outside PET uptake room door |
| T12 | Target | Hallway outside PET/CT exam room door |
| T13 | Target | Janitor closet |
| T14 | Target | Offices and conference room (uncontrolled) |
| T15 | Target | Hallway (uncontrolled) |
| T16 | Target | PET/CT control room |
| T17 | Target | Internal Medicine H above PET uptake room (uncontrolled) |
| T18 | Target | Internal Medicine H above PET WC (uncontrolled) |
| T19 | Target | Internal Medicine H above PET/CT exam room (uncontrolled) |
| *Additional shielding is provided to the technician doing the injection by standard 20 mm lead Pb shield used routinely at Rambam since 2012. **All areas are controlled areas unless stated otherwise. |  |  |

Figure 1: Floor Plan with Reference Codes for Radiation Shielding (See Lists Previous Pages)

Sample calculation: Exposure to the Internal Medicine H wing above the uptake room.

Patients are assumed to be injected with an average of 10 mCi (370 MBq) 18F-FDG. Patients are assumed to be in the injection location for 5 minutes, in one of 3 uptake locations for 50 minutes, and then in the WC (toilet) for 5 minutes. After voiding, the patient is assumed to be scanned for 20 minutes. All 3 uptake positions and the injection position are assumed to occupied and directly under the target. Target occupancy is 1 and maximum threshold is 4 μSv/week for Internal Medicine H (uncontrolled area). The floor to floor distance is 3.38 m. All workers or patients are assumed to be sitting at 0.5 m above the floor. Vertical distance from source to target is then:

d = 3.38 – 0.5 + 0.5 = 3.38 m

The dose rate reduction factor during 5 minutes of injection is:

Rt=[1/ln(2)](T1/2/t){1–exp[–ln(2)t/T1/2]}

RtU1=[1/ln(2)](109.77/5){1–exp[–ln(2)5/109.77]} = 0.98438

The dose rate reduction factor during 50 minutes of uptake is:

RtU2=[1/ln(2)](109.77/50){1–exp[–ln(2)50/109.77]} = 0.85752

Given that 10 mCi is (10 x 37) MBq, the unshielded dose from 1 patient being injected is:

D1 = ΓA0tU1RtU1/d2

D = 0.092(10 x 37)(5/60) 0.98438/3.382 =  0.24442 μSv

Given that 10 mCi is (10 x 37) MBq, the unshielded dose from 1 uptake patients is:

D2 = ΓA0tU2RtU2/d2

D2 = 0.092(10 x 37)(50/60) 0.85752/3.382 =  2.1292 μSv

Given 160 patients per week, the weekly unshielded dose from one uptake patient is

D = 160 x (0.24442+2.1292) = 379.8 μSv

The transmission factor for a 40 cm thick concrete ceiling of the uptake room between the waiting patients and internal medicine above is:

B = {(1+(β/α))eαγx−(β/α)}(−1/γ)

B = {(1+(–0.1161/0.1539))e0.1539(2.0752)(34) – (–0.1161/0.1539)}(−1/2.0752)

B = 0.010503

The shielded weekly dose from 160 patients into Internal Medicine H above the uptake room is:

D = 0.010503 x 379.8 = 3.99 μSv

The dose from these patients through the ceiling into Internal Medicine H will be smaller on average since most of the high energy photons will be striking the ceiling at oblique angles. This both increases the effective distance between the source and the target (reducing the intensity via distance) and increases the transit thickness of the ceiling (reducing the intensity via absorption).

The patient being injected and the uptake patients are separated in calculation to account for geometry. In this case, there was no difference because the shortest distance straight up from patient to target was assumed and it is the same for all patients in the waiting room. In all calculations, the most conservative estimate was used by assuming the shortest distance from patient to target.

Discussion on Calculations found in Tables

Tables with dose calculations are found below. Figure 1 is a suitable guide for the targets listed in the tables.

Table 1

All patients will spend some time in the injection location, the WC (toilet), and the exam room, and these were all taken as sources. Each patient will be in one of 3 locations for the uptake wait, so the most conservative (highest source to target exposure) of these 3 locations was taken to calculate exposure to the target.

The exposure rate in the SPECT/CT control room is under 40 μSv/week (controlled area limit). The value 35.5 μSv/week is calculated assuming that all injected patients are directly opposite the technician with both the source and target within 30 cm of each side of a 36 cm concrete wall. This is a conservative estimate since this minimal distance would be rare.

The exposure rate in the uptake room at the injection location, 25.8 μSv/week, is under 40 μSv/week (controlled area limit) but requires the technician to use a 20 mm Pb body shield. This has been standard practice at Rambam Health Care Campus since August 2012.

Table 2

The exposure rate in the PET control room is under 40 μSv/week (controlled area limit). The value 39.0 μSv/week would be calculated using the weakest shield (assuming 4 mm Pb equivalent glass window) and assuming that the technician is within 30 cm of the window. The control room consoles move the technician further than 30 cm away from the window and the 20 cm thick concrete wall provides more shielding than the window to most of the body. Consequently, this estimate is conservative and the average exposure would be 5.9 μSv/week.

The corridor is in a controlled area used by radiation workers and injected patients. The walls provide sufficient shielding to meet the 40 μSv/week dose limit. However, the area just outside the door remains problematic: the occupancy factor needs to be 1/16 instead of a more conservative 0.25 in order to meet the dose requirements. We believe this is reasonable as the hall is well-monitored by radiation technicians, nurses, and physicists. People staying near the door of an uptake room are always informed that it is an area of high radiation field and that they are required to remove themselves.

Table 3

Outdoor exposure is within the 4 μSv/week (uncontrolled area limit). In fact, this area is almost never used since it is in a corner behind a structural pillar.

Table 4

Exposure in the Internal Medicine H wing above the PETuptake room is under 4 μSv/week (uncontrolled area limit): 3.9 μSv/week. This is a conservative estimate because oblique lines-of-sight cause the radiation flux to pass through more shielding in the ceiling than the thicknesses used here. Also, oblique angles increase the effective distances to larger than the 3.38 m floor-to-floor distance. The calculations are done as if the uptake patients were a single source under the target in the Internal Medicine H wing.

Table 5

Maximum radiation fields were calculated assuming that all locations were filled with injected patients, accounting for decay and shielding. Three example tallies are shown for the highest fields. The highest case is 11.4 μSv /h which is well under the 20 μSv/h limit for public exposure. This is the radiation field outside the doors to the uptake room.

Conclusion

The room design for the new 2019 PET/CT, Rambam given by the plan below provides sufficient shielding to workers and the public.

| Table 1: Controlled Areas Part A: Limit 40 μSv/week (All concrete is 2.35 g/cm3 density) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Target | Source | Distance (m) | Shielding Code | Shielding Thickness (cm) | Shielding Material | Occupancy Factor | Weekly Unshielded Dose (μSv) | Transmission Factor | Weekly Dose (μSv) |
| Technician | S1 | 1 | Shield* | 2 | Lead (Pb) | 0.25 | 111.7 | 0.0535 | 6.0 |
| Injecting | S2 | 2 | Shield | 2 | Lead (Pb) | 0.25 | 235.7 | 0.0535 | 12.6 |
| T1 | S5 | 3.3 | D4 | 0 | None | 0.25 | 7.2 | 1 | 7.2 |
|  |  |  |  |  |  |  |  | Total: | 25.8 |
| SPECT | S1 | 0.96 | W2 | 36 | Concrete | 1 | 484.8 | 0.0077 | 31.6 |
| Control | S2 | 0.96 | W2 | 36 | Concrete | 1 | 4091.8 | 0.0077 | 3.7 |
| Room | S5 | 4.3 | W2 | 36 | Concrete | 1 | 17.1 | 0.0077 | 0.1 |
| T2 |  |  |  |  |  |  |  | Total: | 35.5 |
| Hallway by | S1 | 1.1 | W1 | 16 | Concrete | 0.25 | 92.3 | 0.1662 | 15.3 |
| PET | S2 | 2.9 | W1 | 16 | Concrete | 0.25 | 112.1 | 0.1662 | 18.6 |
| Uptake | S5 | 2.3 | W4 | 36 | Concrete | 0.25 | 14.9 | 0.0143 | 0.2 |
| T9 | S6 | 5.7 | W8 | 0.2 | Lead (Pb) | 0.25 | 6.1 | 0.7873 | 4.8 |
|  |  |  |  |  |  |  |  | Total: | 39.0 |
| Hallway by | S1 | 3.0 | W1 | 16 | Concrete | 0.25 | 12.4 | 0.1662 | 2.1 |
| PET/CT | S2 | 4.7 | W1 | 16 | Concrete | 0.25 | 42.7 | 0.1662 | 7.1 |
| Exam | S5 | 2.8 | W7 | 27 | Concrete | 0.25 | 10.1 | 0.0489 | 0.3 |
| T10 | S6 | 3.3 | W8 | 0.2 | Lead (Pb) | 0.25 | 18.2 | 0.7873 | 14.4 |
|  |  |  |  |  |  |  |  | Total: | 23.8 |
| *This shielding is provided to the technician doing the injection by standard 20 mm lead Pb shield used routinely at Rambam since 2012. |  |  |  |  |  |  |  |  |  |

| Table 2: Controlled Areas Part B: Limit 40 μSv/week (All concrete is 2.35 g/cm3 density) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Target | Source | Distance (m) | Shielding Code | Shielding Thickness (cm) | Shielding Material | Occupancy Factor | Weekly Unshielded Dose (μSv) | Transmission Factor | Weekly Dose (μSv) |
| Hallway by | S1 | 1.5 | D1 | 0.4 | Lead (Pb) | 0.0625 | 12.4 | 0.6021 | 7.5 |
| PET | S2 | 2.9 | D1 | 0.4 | Lead (Pb) | 0.0625 | 28.0 | 0.6021 | 16.9 |
| Uptake | S5 | 2.3 | W4 | 32 | Concrete | 0.0625 | 3.7 | 0.0143 | 0.1 |
| Door | S6 | 5.7 | W8 | 0.2 | Lead (Pb) | 0.0625 | 1.5 | 0.7873 | 1.2 |
| T11 |  |  |  |  |  |  |  | Total: | 25.6 |
| Hallway by | S1 | 3.0 | W1 | 16 | Concrete | 0.0625 | 3.1 | 0.1662 | 0.5 |
| PET/CT | S2 | 4.7 | W1 | 16 | Concrete | 0.0625 | 10.74 | 0.1662 | 1.8 |
| Exam | S5 | 4.8 | W7 | 27 | Concrete | 0.0625 | 0.9 | 0.0486 | 0.0 |
| Door | S6 | 2.8 | D2 | 0.2 | Lead (Pb) | 0.0625 | 6.3 | 0.7873 | 5.0 |
| T12 |  |  |  |  |  |  |  | Total: | 7.3 |
| Janitor | S1 | 3.0 | D1 | 0.4 | Lead (Pb) | 0.0625 | 26.2 | 0.6021 | 15.8 |
| Closet | S2 | 4.7 | D1 | 0.4 | Lead (Pb) | 0.0625 | 10.7 | 0.6021 | 6.4 |
| T13 | S5 | 3.5 | W7 | 27 | Concrete | 0.0625 | 19.2 | 0.0308 | 0.6 |
|  | S6 | 2.8 | W8 | 0.2 | Lead (Pb) | 0.0625 | 6.3 | 0.7873 | 5.0 |
|  |  |  |  |  |  |  |  | Total: | 27.8 |
| PET/CT | S6 | 3.5 | W11 | 20 | Concrete | 1 | 64.8 | 0.0904 | 5.9 |
| Control |  |  |  |  |  |  |  | Total: | 5.9 |
| Room |  |  |  |  |  |  |  |  |  |
| T16 |  |  |  |  |  |  |  |  |  |

| Table 3: Uncontrolled Areas Part A: Limit 4 μSv/week (All concrete is 2.35 g/cm3 density) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Target | Source | Distance (m) | Shielding Code | Shielding Thickness (cm) | Shielding Material | Occupancy Factor | Weekly Unshielded Dose (μSv) | Transmission Factor | Weekly Dose (μSv) |
| Outside | S1 | 4.2 | W3 | 31 | Concrete | 0.0625 | 1.6 | 0.0167 | 0.0 |
| behind | S3 | 1.0 | W3 | 31 | Concrete | 0.0625 | 37.7 | 0.0167 | 3.9 |
| Pillar |  |  |  |  |  |  |  | Total: | 4.0 |
| T3 |  |  |  |  |  |  |  |  |  |
| Stairwell | S1 | 4.9 | W3 | 31 | Concrete | 0.0625 | 1.2 | 0.0167 | 0.0 |
| T4 | S4 | 1.2 | W3 | 31 | Concrete | 0.0625 | 163.7 | 0.0167 | 2.7 |
|  | S5 | 4.3 | W5 | 20 | Concrete | 0.0625 | 1.1 | 0.1227 | 0.1 |
|  |  |  |  |  |  |  |  | Total: | 2.8 |
| Janitor | S1 | 3.8 | W4 | 32 | Concrete | 0.0625 | 1.9 | 0.0143 | 0.0 |
| Closet | S4 | 1.0 | W4 | 32 | Concrete | 0.0625 | 235.7 | 0.0143 | 3.4 |
| T5 | S5 | 3.3 | W5 | 20 | Concrete | 0.0625 | 1.8 | 0.0904 | 0.2 |
|  |  |  |  |  |  |  |  | Total: | 3.6 |
| WC | S1 | 3.6 | W4 | 32 | Concrete | 0.0625 | 2.2 | 0.0143 | 0.0 |
| T6 | S4 | 1.7 | W4 | 32 | Concrete | 0.0625 | 81.6 | 0.0143 | 1.2 |
|  | S5 | 0.8 | W5 | 20 | Concrete | 0.0625 | 30.8 | 0.0904 | 2.8 |
|  |  |  |  |  |  |  |  | Total: | 4.0 |

| Table 4: Uncontrolled Areas Part B: Limit 4 μSv/week (All concrete is 2.35 g/cm3 density) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Target | Source | Distance (m) | Shielding Code | Shielding Thickness (cm) | Shielding Material | Occupancy Factor | Weekly Unshielded Dose (μSv) | Transmission Factor | Weekly Dose (μSv) |
| WC | S5 | 0.8 | W6 | 20 | Concrete | 0.0625 | 30.8 | 0.0904 | 2.8 |
| T7 |  |  |  |  |  |  |  | Total: | 2.8 |
| Hallway | S5 | 0.8 | W7 | 27 | Concrete | 0.25 | 123.3 | 0.0308 | 3.8 |
| T8 |  |  |  |  |  |  |  | Total: | 3.8 |
| Offices | S6 | 3.3 | W9 | 25 | Concrete | 1 | 72.9 | 0.0419 | 3.1 |
| T14 |  |  |  |  |  |  |  | Total: | 3.1 |
| Hallway | S6 | 3.3 | W10 | 20 | Concrete | 0.25 | 18.2 | 0.0904 | 1.6 |
| T15 |  |  |  |  |  |  |  | Total: | 1.6 |
| Above | S1 | 3.38 | C1 | 34 | Concrete | 1 | 39.1 | 0.0105 | 0.4 |
| Uptake | S2 | 3.38 | C1 | 34 | Concrete | 1 | 330.1 | 0.0105 | 3.5 |
| T17 |  |  |  |  |  |  |  | Total: | 3.9 |
| Above | S5 | 3.38 | C2 | 17 | Concrete | 1 | 27.6 | 0.1428 | 3.9 |
| PET WC |  |  |  |  |  |  |  | Total: | 3.9 |
| T18 |  |  |  |  |  |  |  |  |  |
| Above | S6 | 2.88 | C3 | 25 | Concrete | 1 | 95.7 | 0.0419 | 4.0 |
| PET WC |  |  |  |  |  |  |  | Total: | 4.0 |
| T19 |  |  |  |  |  |  |  |  |  |

| Table 5: Radiation Fields Assuming Fully Injected Patients at All Locations (370 MBq) with Decay (20 μSv/h max.) |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Target | Source | Distance (m) | Shielding Code | Shielding Thickness (cm) | Shielding Material | Unshielded Radiation Field (μSv/h) | Transmission Factor | Shielded Radiation Field (μSv/h) |
| Controlled | S1 | 1.5 | D1 | 0.4 | Pb | 10.69 | 0.6021 | 6.44 |
| Hallway | S2 | 2.9 | D1 | 0.4 | Pb | 3.92 | 0.6021 | 2.36 |
| Outside | S3 | 4.5 | D1 | 0.4 | Pb | 1.63 | 0.6021 | 0.98 |
| Uptake | S4 | 4.5 | D1 | 0.4 | Pb | 1.63 | 0.6021 | 0.98 |
| T11 | S5 | 2.3 | W4 | 32 | Concrete | 4.41 | 0.0143 | 0.06 |
|  | S6 | 5.7 | W8 | 0.2 | Pb | 0.71 | 0.7873 | 0.56 |
|  |  |  |  |  |  |  | Total: | 11.38 |
| Public | S1 | 3.6 | W4 | 32 | Concrete | 1.86 | 0.0143 | 0.03 |
| WC | S2 | 3.6 | W4 | 32 | Concrete | 2.54 | 0.0143 | 0.04 |
| T6 | S3 | 3.8 | W4 | 32 | Concrete | 2.28 | 0.0143 | 0.03 |
|  | S4 | 1.7 | W4 | 32 | Concrete | 11.41 | 0.0143 | 0.16 |
|  | S5 | 0.8 | W5 | 20 | Concrete | 36.41 | 0.0904 | 3.29 |
|  |  |  |  |  |  |  | Total: | 3.55 |
| Public | S5 | 0.8 | W6 | 20 | Concrete | 36.41 | 0.0904 | 3.29 |
| WC |  |  |  |  |  |  | Total: | 3.29 |
| T7 |  |  |  |  |  |  |  |  |